cold sterilization of medical devices and materials by plasma DC glow discharge

A.V.Khomich1, I.A.Soloshenko1, V.V.Tsiolko1, I.L.Mikhno2

1Institute of Physics NAS Ukraine, pr.Nauki 46, 252022, Kiev, Ukraine

2Institute of Epidemiology and Infective Diseases, sp.Protasov Yar 4, 252038, Kiev, Ukraine

ABSTRACT

This report presents the results of investigation the efficiency of sterilization of medical devices directly in the plasma of DC glow discharge. Sterilization regime is independent on the discharge chamber loading by processed instruments. It is shown that effective sterilization is reached using either air or other gases (argon, nitrogen, oxygen) with comparatively low power density (0.003 - 0.03 W/cm3). Sterilization time comprises less than 10 minutes. Temperature of the processed instruments does not exceed 650C. Usage of this discharge results in essential simplifying of the sterilizer, lowering its cost and increasing its safety for servicing personnel and also improvement of its operational characteristics in comparison with the devices utilizing RF or VHF discharges.

1  INTRODUCTION

A new technique for sterilization of medical articles and instruments by gas discharge plasma is now developed intensively. This technique possesses a number of advantages as compared to known dry heat, steam autoclave and gas (EtO) techniques. Principal advantages are low heating of articles and short time of sterilization. However, a number of drawbacks is inherent to existing sterilizers as follows: 1) Usually plasma in such devices is generated by RF or VHF discharge [1-2]. It requires, at first, the usage of RF or VHF generators, which increases complexity and cost of the device. At second, radiation of the generators creates probable danger for servicing personnel and thus extra shielding is required. 2) In a majority of designs the sterilization occurs in the plasma spread from the discharge region by diffusion [1-2] rather than immediately in the discharge. It increases the losses of active particles at internal elements of the sterilizer chamber and thus increases the sterilization time. 3) In the devices providing immediate the sterilization immediately in the discharge, placement of processed instruments affects the discharge regime essentially [3].

This proceeding is devoted to studying the efficiency of sterilization of medical articles immediately in the gas discharge of direct current with special configuration of the electrodes so that the chamber loading by processed instruments does not affect essentially on the discharge glowing and, consequently, on the sterilization regime.

2  TECHNICAL DETAILES

2.1  Experimental technique

Investigations have been performed at the plasma sterilizer setup with chamber volume of 25 liters, which was evacuated by a forepump down to residual pressure of 1.3 Pa. Air, argon, nitrogen and their mixtures were used as operating gas. Pressure of the gas was established by means of feeding valve in a range of 8  27 Pa. Density of power introduced into the discharge was varied in a range 0.003  0.03 W/cm3. Measurements of the principal plasma parameters (concentration of charged particles and electron energy distribution function) were performed by single and dual Langmuire probes. Since medical practice dictates strict requirement of avoiding damage of articles made of thermolabile plastics by their overheating, especial attention has been paid to the measurements of articles temperature in various sterilization regimes. Here it should be noted that performing temperature measurements by thermo-pair placed into the discharge (as it was used in [1]) results in significant error. In this proceeding the thermo-pair was used for direct measurements of the temperature of sterilized articles. It has been shown that the temperature grows up with time flow for practically all discharge regimes, however, it does not exceed 650C during a time of complete sterilization.

Metallic and glass Petri dishes with internal surface of 10 cm2 were used as test objects. Strains of microorganisms in vegetative form (E. Coli, St. Aureus), spore form (Bac. Subtilies), as well as hepatitis B virus were used in the experiments. The majority of studies has been performed using Bac. Subtilies as the strain, which is most resistant to sterilizing action of the plasma. Suspension of spores Bac. Subtilies containing not less than 95 % of spores was used for infection of the test objects. Initial population of spores on test objects was varied in a range of 105  109. During the preparation of the samples, spores suspension was deposited homogeneously onto internal surface of Petri dishes, which in the consequent were dried at temperature of 370C. 24 hours later of the sterilization accomplishing the direct seeding was performed onto rigid (liquid, if required) nutrient media. Meat-peptone agar was used as rigid nutrient medium, and meat-peptone bouillon or thioglicole medium were used as liquid ones. Test samples were incubated for 48 hours and a colony count was then performed.

2.2  Results and discussion

The experiments have demonstrated that the sterilization time decreases essentially with the increase of plasma density. For this reason essential non-homogeneity of the plasma must be avoided in commercial sterilizers (sterilization time for the set of instruments is determined by one actual for instrument placed in a region with minimal plasma density). In the discharge under study satisfactory degree of the plasma homogeneity (less than 25 %) for the major portion of chamber volume was observed in a pressure range of 8  13 Pa for all used gases. At a pressure of 27 Pa the non-homogeneity might reach 50 %. An example of spatial distribution of the plasma density in the plasma sterilizer chamber is given in Fig. 1. One can see that the region with weak non-homogeneity of the plasma embraces more than 80 % of the chamber volume. Chamber loading by the instruments to be sterilized does not affect essentially the regime of discharge glowing and the plasma homogeneity.

Fig. 1. Spatial distribution of the plasma density in the chamber of plasma sterilizer. p = 10 Pa, air, power density 0.008 W/cc.

Fig. 3: Kill curves for Bac.Subtilis obtained by plate count technique. Initial population 107, power density 0.01 W/cc, air.

Fig. 2: Kill curves for Bac.Subtilis obtained by plate count technique. Power density 0.01 W/cc, p = 10.6 Pa, air.

For the demonstration possibilities of developed sterilizer and of the plasma sterilization technique in general we would like to show certain results of medical and biological studies. Fig. 2 exhibits kill curves for two initial spores populations (107 and 108). Each point in the graph represents averaged results of counting for 6 test samples. One can see from this figure that, at first, both curves of survival possess specific feature - sharp falling down (6-7 orders of magnitude) during initial minutes of processing, and than essentially more smooth decrease in remaining time. At second, one can see distinctly, that increase of spores population from 107 to 108 results in essential growth of sterilization time. Reason for such behavior may consist in clumping and occlusion of spores on the surface of test samples, which result in a delay of penetrating the sterilizing agents to the spores and, consequently, in extending the survival time. Another evidence for this assumption is given by the fact that in a range of spores concentrations 104  106 sp/cm2 the sterilization time does not depend on spores concentration value, and the curves possess practically exponential behavior. Obviously such behavior is common for various designs of plasma sterilizers, and the concentration of 107 sp/cm2 is an upper limit for one acceptable for the plasma sterilization technique.

The studies of influence of air pressure in the chamber and specific power introduced into the discharge on the sterilization process. One can see from the Figs. 3,4 that the time of reaching 10-6 sterility assurance level is practically independent on air pressure varied in a range of 10  26 Pa and on specific power variations in a range of 0.003  0.01 W/cm3. The investigations have shown as well that the increase of specific power more than 0.01 W/cm3 is not reasonable since it does not result in essential decrease of the sterilization time. We’d like to note, that plasma concentration comprises 2109 cm-3 for this value of specific power.

Fig. 4. Kill curves for Bac.Subtilis obtained by plate count technique. Initial population 107, p = 10.6 Pa, air.

Fig. 5: Kill curves for Bac.Subtilis obtained by plate count technique. Initial population 107, p = 26.6 Pa, power density 0.01 W/cc.

Commonly it is suggested to use complex mixtures of operating gases in the plasma sterilizers, that results in the increase of device complexity and cost. We have accomplished the experiments on determining the sterilization efficiency for the cases of various operating gases. Fig. 5 shows the kill curves for usage of either argon or air as operating gas. One can see that air plasma provides stronger sterilizing ability. On the contrary, usage of oxygen results in the sterilization time decrease just in 1.5 - 2 times. Thus the conclusion can be formulated that usage of atmospheric air as operating gas is the most reasonable way for commercial device.

The essential aspect of determining the optimal conditions for the plasma sterilization consists in establishing of relative contributions of principal sterilizing factors:

a)	chemically active electrically neutral plasma particles (radicals, excited atoms and molecules);

b)	slow ions and electrons of the plasma;

c)	fast ions and electrons of the plasma;

d)	ultraviolet radiation of the plasma;

For determining the roles performed by charged particles in the sterilization process the experiments have been accomplished with various potentials applied to the Petri dishes (particularly, potentials of discharge cathode, discharge anode and floating potential of the plasma). In these three cases currents and energies of both electrons and ions of the plasma, targeting the dishes, are essentially different. Would the role of charged particles be essential, the sterilization times in mentioned regimes should be different. The experiments, however, have demonstrated that in all three regimes the curves of survival are identical. It is illustrated in Fig. 6 for two potentials of the dishes: floating potential and anode one.

Resuming we can state that principal role in the plasma sterilization is played by active electrically neutral plasma particles and ultraviolet radiation of the plasma.

Fig. 6: Kill curves for Bac.Subtilis obtained by plate count technique. Initial population 108, power density 0.09 W/cc, p = 10.6 Pa, air.

3  CONCLUSIONS

Thus the following issues are established in  current proceeding:

1.	Glowing discharge on direct current is effective one for sterilization of medical instruments. It does not require RF or VHF generator, provides short sterilization time, ensures low temperature of the processed instruments, possesses high efficiency of energy usage and does not change the glowing regime when the chamber is loaded by processed instruments.

2.	Maximal allowed spores concentration in the plasma sterilization technique should not exceed 107 spores/cm2.

3.	Atmospheric air is the most preferable gas to be used as operating medium in the sterilizer.

4.	The principal sterilizing agents of the plasma are ultraviolet radiation and active electrically neutral particles.

4  ACKNOLEDGMENTS

This work was supported in part by Anti-AIDS National Committee of Ukraine and Science and Technology Center in Ukraine under contract # 57.

5  REFERENCES

[1]	A.Jacob, US patent 5413760.

[2]	B.A.Cambell, K.A.Moulton, US patent 5413760.

[3]	K.Otsuka, Japanese patent 60-58662.